## Slide 1
C. Graf (a), M. Langlois (a), M. Loupias (a), E. Thiébaut (a), M. Tallon (a), J. Degallaix (b)
(a) Centre de Recherche Astrophysique de Lyon (CRAL), avenue Charles-André - Saint-Genis Laval
(b) Laboratoire des Matériaux Avancés (LMA), 7 Avenue Pierre de Coubertin 69622 - VILLEURBANNE Cedex
Integrated Mach-Zehnder (iMZ) wavefront sensor
Wavefront sensing at nanometric scale for XAO and high precision metrology
WFS Workshop in the VLT/ELT era
Porto, 19-21 October 2022

## Slide 2
2 glass plates of exact same width
λ/4 beam splitter layer (rs, ts)
reflective dot : pinhole ( ⌀ ≈ λ/D )
Creation of the reference signal
iMZ = self-referenced interferometer
spatial filtering
iMZ schematic representation
The iMZ :
integrated Mach-Zehnder
λ/8 layer + reflective layer (rm)
λ/8
a ≈ λ/D
Mach-Zehnder interferometer with glass propagation
Fourier-filtering family of wavefront-sensor
Notes: -Dire au début que c'est un interferometre classique
operating range : 600 nm  900 nm (new prototype  goes 850 um)
Ti305 beam splitter layer (lambda/4)
50 um diameter pinhole
lambda/8 OPD of SiO2
(OPTOPRIM)

## Slide 3
IA
IS
Two complementary outputs IS & IA
Symmetric output
Anti-symmetric output
, : calibration coefficients, depending on the coatings and of the filtering ()
: incoming photon number
: intensity distribution on the pupil
: incoming phase
: reference phase
The iMZ :
integrated Mach-Zehnder
a ≈ λ/D
λ/8
I
≈30 %
≈70 %
iMZ schematic representation

## Slide 4
The iMZ :
integrated Mach-Zehnder
λ/8
IA
IS
I
Two complementary outputs IS & IA
Symmetric output
Anti-symmetric output
Complementary information
Sum of the 2 outputs : separate retrieval of phase and intensity
Signal linear with small phases
iMZ schematic representation
≈30 %
≈70 %
a ≈ λ/D

## Slide 5
The XAO test bench at CRAL :
iMZ
Camera Neo
DM
12x12 actuators
3.5 microns stroke
SLM
256x256 pixels
4 nm precision
Camera Zyla

## Slide 6
Calibration method :
Method : linear inverse problem resolution
Objective : extract the calibration phase and the coefficient , from iMZ signal for each pixel of the pupil
Hypothesis : - perfect filtering : -
- , are phase independent
Calibration model (anti-symmetric output, = ) :
Linearization of the model : introduction of a set of known modulations with a deformable mirror (DM) :
where = and
Estimator (Least Mean Square Estimator) :
with the data, the direct model matrix, the parameters vector and the noise matrix such as
Independent calibration of the iMZ coefficients per subaperture using inverse approach
Optimized to accurately disentangle phase and intensity

## Slide 7
Phase extraction :
Extraction of the phase using the calibration coefficients and reversing the model :
Method captures non uniform pupil intensity
Non uniformity coming from the DM
Ratio ≈

## Slide 8
Simulation results : Increasing dynamic
iMZ signal without unwrapping strategy
Modulation : introduction of a known phase step and its inverse
MZ in
Unwrapping
Residual 1st iteration
(rad2)
Phase (radians)
Residual 2st iteration
(rad2)
(radians)
1D. C. Ghiglia and L. A. Romero, “Robust two-dimensional weighted and unweighted phase unwrapping that uses fast transforms and iterative methods,” J. Opt. Soc. Am. A 11, pp. 107–117, Jan 1994.
Unwrapping1 on the modulated signal  up to several microns
M. J. Townson, O. J. D. Farley, G. O. de Xivry, J. Osborn, and A. P. Reeves, “AOtools: a python package for adaptive optics modelling and analysis,” Optics Express 27, p. 31316, oct 2019.

## Slide 9
Noise propagation :
Low noise propagation close to the theoretical limit of non modulated Pyramid WFS
O. Guyon, “Limits of adaptive optics for high-contrast imaging,” The Astrophysical Journal 629, pp. 592–614, aug 2005
Theoretical limit :
(photon noise only)

## Slide 10
Simulation results : Wavefront reconstruction
Figure 2 : MZ phase reconstruction. In red entrance phase, in green iMZ output phase, in orange phase after modulation and in blue after unwrapping.
-
-
Noise free simulation :
D = 8m, 40x40 subapertures, medium seeing
Accurate wavefront reconstruction in one iteration (SR > 90%)
modulation + unwrapping
small phase regime
Perfect wavefront reconstruction in the second iteration (SR > 99%)
phase input
imZ phase output
residual 1st iteration
residual 2nd iteration

## Slide 11
Simulation results : Wavefront reconstruction
Figure 2 : MZ phase reconstruction. In red entrance phase, in green iMZ output phase, in orange phase after modulation and in blue after unwrapping.
-
-
Realistic flux simulation : mv = 9, framerate = 500 Hz
D = 8m, 40x40 subapertures, medium seeing
SR ≈ 75 % for bootstrapping (modulation + unwrapping)
phase input
imZ phase output
residual 1st iteration
residual 2nd iteration
modulation + unwrapping
small phase regime
SR ≈ 96 % for small phases reconstruction

## Slide 12
Experimental results :
Comparison of the phase from the two outputs :
residual variance: rad2
Phase step introduced by a SLM (≈150 nm)
Anti-symmetric output
Symmetric output
Difference
iMZ high sensitivity to phase steps + measure of amplitude

## Slide 13
Experimental results :
Correction loop : phase screen on a 256x256 SLM correction with a 12x12 DM
Accurate measurement at 1st iteration
Strehl ratio > 90% after 2 iterations
raw pupil
phase pupil
image
radians
ADU
ADU
PSF on the coronographic path
(behind the pinhole)

## Slide 14
Experimental results :
Flux calibration : necessity to take into account the flux variation on the pinhole due to the phase diversity
Refining the calibration model :
PSF image behind the pinhole :
(reference phase)
( random phase applied)
with
= Number of reflected photons on the pinhole for phase
without flux ajustement
with flux ajustement

## Slide 15
Take-away :
Accurate measurement of both phase (at nanometer scale) and amplitude of the signal (scintillation)
Efficient calibration method to extract the phase from the iMZ signal which captures the non uniformity of the intensity in the pupil
Good sensitivity : low noise propagation
Simulation tool to realistically simulate the iMZ response to different phase : accurate phase reconstruction outside the iMZ linearity range thanks to modulation + unwrapping strategies
iMZ very efficient for XAO without first stage correction (except tip-tilt)